**Centura Health Logger Configuration**

**For LogTag TRED30-7R Logger**

* Configure Loggers all loggers for:
	+ Push Button Start
	+ Record Reading every 30 min
	+ Begin recording after a delay of 5 min
	+ No Password
	+ Advance Options:
		- Select alarm remains on even if readings return to non-violation range
		- Temperature display unit in Fahrenheit
		- Select Show total Summary Days Collected
		- Allow logging stop with STOP button
		- Enable Power Save – turn off display after 30 seconds
* Configure Alarms:
	+ Fridge – Set high alarm to 46.1°F (8.1°C) & low to 35.9°F (1.9°C)
	+ Freezer – Set high alarm to 5.1°F (-14.9°C) & low to -40°F (-40°C)